## Slide 1: DICOM WG 30 Small Animal Imaging DICOM Training Introduction and Basic Concepts
David Clunie (dclunie@dclunie.com)
PixelMed Publishing

## Slide 2: DICOM – Learning Objectives
Brief history
File formats and Storage Services
Data Set encoding principles
Transfer Syntaxes
Information Object Definitions
Storage SOP Classes
Other DICOM services than storage
Conformance statements
Association Negotiation
DICOM Message Service Elements (DIMSE)
DICOM Upper Layer Protocol (DUL)

## Slide 3: Who Needs to Know What?
User/customer
  services and objects relevant to domain (esp. storage)
  conformance statement interpretation and matching
Installer/integrator
  + network addressing, more detail about services
  +/- integration with non-imaging (IS, workflow) systems
Application designer/developer
  detailed knowledge of (relevant) services, objects, attributes
  uses toolkit to abstract (most) encoding/network details
Problem solver/troubleshooter
  + data sets: interpret dumps and validation tool output
  + network: DIMSE/DUL to interpret network logs/traces/dumps
Toolkit designer/developer
  everything: objects, encoding, services, network
Standard writer: new modality
  IOD structure, reusable modules, encoding limits, precedents

## Slide 4: DICOM – Brief History
1982 – 1st PACS Conference – session on standards
1982 – AAPM Report 10 – Standard Format for Image Interchange
1983 – ad hoc meeting between FDA, ACR & NEMA
1983 – 1st meeting of ACR-NEMA “Digital Imaging and Communications Standards” Committee
1985 – ACR-NEMA 300-1985 (“version 1.0”) issued
1988 – ACR-NEMA 300-1988 (“version 2.0”) issued
1990 – Inter-vendor testing of version 2.0 at Georgetown
1992 – Trial of DICOM (“version 3.0) at RSNA
1993 – DICOM 3.0 issued

## Slide 5: DICOM – Brief History
1982 – 1st PACS Conference – session on standards
1982 – AAPM Report 10 – Standard Format for Image Interchange
1983 – ad hoc meeting between FDA, ACR & NEMA
1983 – 1st meeting of ACR-NEMA “Digital Imaging and Communications Standards” Committee
1985 – ACR-NEMA 300-1985 (“version 1.0”) issued
1985 – IEEE 802.3 Ethernet (based on 1976 Metcalfe)
1986 – Aldus TIFF (version 3; prior versions drafts only)
1987 – CompuServe GIF
1988 – ACR-NEMA 300-1988 (“version 2.0”) issued
1990 – Inter-vendor testing of version 2.0 at Georgetown
1992 – Trial of DICOM (“version 3.0) at RSNA
1992 – JPEG (ITU T.81; ISO 10918-1 1994)
1993 – DICOM 3.0 issued

## Slide 6: DICOM – Brief History
ACR-NEMA versions 1 and 2
  50-pin 16 bit parallel interface
  no network (assumed “network interface unit”)
  layered
  messages with commands and data
  tag-value pairs
  described patients, studies, images
  described modality, acquisition, 3D position, etc.

## Slide 12: DICOM – Brief History
ACR-NEMA versions 1 and 2
  50-pin 16 bit parallel interface
  no network (assumed “network interface unit”)
  layered
  messages with commands and data
  tag-value pairs
  described patients, studies, images
  described modality, acquisition, 3D position, etc.
DICOM “3.0”
  TCP/IP network protocol (and OSI semantics)
  “object-oriented” description & conformance

## Slide 13: DICOM – Scope
Implicit in the standard’s name: “Digital Image and Communications in Medicine”
Images and image-related information
All acquisition modalities
  1993: CR, CT, MR, US, NM, SC (DF and DV)
  2014: XA, XRF, DX, MG (+DBT), IO, PT, VL (Photo, Endo, Slide), WSI, OCT (OP,IV)
  RT, SR (+CAD), waveforms, OP measurements and maps, surface scans, implants, segmentations
Storage + other services

## Slide 14: DICOM – Scope
Interchange of medical images
  radiology, cardiology, pathology .... any ‘ology
  network (TCP/IP) or media (CD, DVD, MOD)
Printing
  grayscale and color
  network rather than point-to-point (TCP/IP)
Information System Integration
  reports, worklists, performed procedures

## Slide 15: Why does DICOM exist ?
Previous solutions closed (proprietary)
Real world requires open standards
  heterogeneous modalities
  heterogeneous workstations
  heterogeneous archives, PACS, etc.
Consortium of
  industry (NEMA)
  users (ACR, ACC, CAP ...)

## Slide 16: Why interchange images ?
Archiving
Interpretation
  softcopy reading
  filmless operation
Advanced processing
  3D visualization, quantitative analysis
  fusion of images from multiple modalities
Teaching, research ...

## Slide 17: Why (open) standards at all ?
Too many permutations of ...
  Acquisition device vendors
    general
      GE, Siemens, Philips, Toshiba, Shimadzu, Hitachi, ...
    special
      Bruker, iTheraMedical, Mediso, PerkinElmer ...
  Processing and archive device vendors
    general
      GE, Siemens, Philips, ... (+ open source)
    special
      InViCRO, ...

## Slide 18: Why not existing standards ?
Didn’t “exist” when ACR-NEMA formed
Pure imaging standards (TIFF, etc.)
  limited support for medical image types
  don’t encode domain specific information
Other domains inappropriate
  military, remote sensing, astronomical, etc.
ISO standards (e.g., IPI) never adopted
Medical standards that didn’t do images
  HL7, MIB (IEEE 1073), etc.
Needed more services than just storage
  query/retrieve, printing, workflow, etc.

## Slide 19: DICOM – Why?
Domain-specific fields
  patient identification & characteristics
  device identification & characteristics
  acquisition protocol parameters
  study management identifiers, dates and times
Need to handle binary bulk pixel data
  e.g., HL7 text based and no image support
Need to be extensible
  evolving technology, private extensions
  variable rather than “fixed” length headers

## Slide 20: Key goals of DICOM
Support interoperability
  NOT (necessarily) interfunctionality
  WITHOUT defining (restricting) architecture
Define conformance
  specific services and objects
  documentation (Conformance Statement)
  negotiation
Consensus standard
Voluntary compliance
Open (license fee free)

## Slide 21: DICOM does NOT define:
PACS or IM Architecture
  Picture Archiving Communications System
  Image Management
Distributed Object Management
Radiology/Hospital Information System
Complete Electronic Medical Record
Integrating the Health Care Enterprise (IHE) does define some aspects of architecture on top of DICOM, HL7, etc.

## Slide 22: What is Interoperability ?
Analogy of web server/browser:
  Interconnectivity - both talk TCP/IP
  Interoperability - both talk HTTP and HTML
  Interfunctionality – approached, not guaranteed:
    “versions” of HTML poorly controlled
    layout and other behavior not constrained by HTML
    optional/additional “features”
    use of extensions (plug-ins, applets, scripts, stylesheets)
Not (always) sufficient for clinical needs
  interoperability is necessary but not sufficient
  reality check on user/customer expectations

## Slide 23: DICOM and Interoperability
For example, conformance to DICOM
  will guarantee network connection
  will guarantee storage of MR image:
    from Modality to Workstation
  will NOT guarantee (but will facilitate):
    Workstation will display image “correctly”
    Workstation can perform analysis (e.g., diffusion tractography)
  facilitated by mandatory attributes for:
    identification, annotation, positioning, etc.
IHE profiles define functionality, e.g., DIFF, PERF, MAMMO, Cardiac NM

## Slide 24: DICOM and Interoperability
“Object-oriented” definition
  data structures, e.g., MR image object
    composite model of real world entities
      patient, study, series
      general image, specialized to MR image
  services, e.g., image storage
  together -> service/object pairs (SOP)
Roles (user or provider) (SCU or SCP)
Role + SOP Class -> Conformance

## Slide 25: DICOM SOP Classes/Roles
MR scanner may say:
  “I am an MR Image Storage Service Class User (SCU)”
Workstation may say:
  “I am an MR Image Storage Service Class Provider (SCP) (amongst other things)” MR images may be transferred

## Slide 26: DICOM SOP Classes/Roles
Angiography device may say:
  “I am an XA Image Storage Service Class User (SCU)”
Workstation may say:
  “I am not an XA Image Storage Service Class Provider (SCP) (though I do support other kinds of images like CT and MR)” This pair cannot transfer XA images

## Slide 27: Why is DICOM so specific ?
For example,
  MR Image
    single frame, 12-16 bit grayscale image
    MR acquisition - pulse sequence parameters
    3D patient relative co-ordinate/vector position
  X-Ray Angiography Image
    multi-frame, 8-10 bit grayscale image
    XA acquisition - radiation/collimation/motion
    dynamic C-arm/table relative positioning

## Slide 28: DICOM SOP Classes/Roles
Workstation may say:
  “I am a Basic Grayscale Print Management Meta SOP Class SCU”
Printer may say:
  “I am a Basic Grayscale Print Management Meta SOP Class SCP” Images may be printed

## Slide 29: DICOM SOP Classes/Roles
Ultrasound scanner may say:
  “I am a Basic Color Print Management Meta SOP Class SCU”
Printer may say:
  “I am only a Basic Grayscale Print Management Meta SOP Class SCP” This pair cannot print images

## Slide 30: Modality is Important
Different acquisition characteristics
  X-Ray-based v. not
  cross-sectional v. projectional (planar)
  transmissive v. emissive
  structural v. functional
  different mechanisms of contrast (endogenous or exogenous)
  hybrids

## Slide 31: Modality is Important
Different image characteristics
  single or multiple channels (grayscale, color)
  use of pseudo-color (palette color)
  broader dynamic range than displayed
  values may be > 8 bits in depth and signed
  non-linear transformation (or linear window)
  physical significance of pixel values (counts, HU, velocity)
  physical significance of size (measurements)

## Slide 32: Modality-specific DICOM IODs
Information Object Definition (IOD)
  one (or more) for each modality
  more if conformance reason for variants
“Composite”
  multiple real-world entities in same IOD
“Object-oriented” (sort of)
  “inherit” same (composite) real-world/information “model”
    patient/study/procedure step/series/image/frame
  re-use common general Modules
    e.g., Patient, General Study, General Image Module
  re-use common technology Modules
    e.g., Contrast/Bolus, Frame of Reference, X-Ray Grid Module
  “specialize” with modality-specific Modules
    e.g., CT Image, MR Image Module

## Slide 33: Composite Information Model

## Slide 34: MR Image IOD

## Slide 35: Image Plane Module
Type: 1 – required, 2 – may be empty if unknown, 3 – optional

## Slide 36: Encoding – Data Elements
Flat list of sorted unique “Data Elements”
  IOD/Module structure is NOT encoded
  Module Attributes encoded as Data Elements
Each Data Element is a Tag-Value pair
  Tag-Type-Value triple with explicit type (VR) encoding
Data Element Tag: pair of 16 bit numbers
  16 bit “group number” (historical distinction)
  16 bit “element number”
  encoded in binary, described in hexadecimal
  e.g., (0008,0020) means “Study Date”
  “meaning” is not encoded – need “data dictionary” to look them up (PS3.6) – i.e., not “self-describing”

## Slide 37: Encoding – Data Elements
Value
  Value Representation (VR) (“type”)
  Value Length (16 or 32 bits; limits max size)
  Value (even length padded; 16 bit parallel legacy)
VRs (“types”)
  binary: US, SS, UL, SL, FL, FD, AT
  bulk binary: OB, OW, OF, OD, UN (“unknown”)
  numeric string: AS, DS, IS, DA, DT, TM
  string: AE, CS, LO, PN, SH, UI
  text: LT, ST, UT

## Slide 39: Example MR Image Dataset
(0x0008,0x0005) CS Specific Character Set VR=<CS> VL=<0x000a> <ISO_IR 100>
(0x0008,0x0008) CS Image Type VR=<CS> VL=<0x0010> <ORIGINAL\PRIMARY>
(0x0008,0x0016) UI SOP Class UID VR=<UI> VL=<0x001a> <1.2.840.10008.5.1.4.1.1.4>
(0x0008,0x0018) UI SOP Instance UID VR=<UI> VL=<0x002e>
<1.2.840.113619.2.1.2.1909421756.1.1.602501582>
(0x0008,0x0020) DA Study Date VR=<DA> VL=<0x0008> <19890203>
(0x0008,0x0021) DA Series Date VR=<DA> VL=<0x0008> <19890203>
(0x0008,0x0022) DA Acquisition Date VR=<DA> VL=<0x0008> <19890203>
(0x0008,0x0023) DA Image Date VR=<DA> VL=<0x0008> <19890203>
(0x0008,0x0030) TM Study Time VR=<TM> VL=<0x0006> <092618>
(0x0008,0x0031) TM Series Time VR=<TM> VL=<0x0006> <093221>
(0x0008,0x0032) TM Acquisition Time VR=<TM> VL=<0x0006> <093302>
(0x0008,0x0033) TM Image Time VR=<TM> VL=<0x0006> <093302>
(0x0008,0x0050) SH Accession Number VR=<SH> VL=<0x0000> []
(0x0008,0x0060) CS Modality VR=<CS> VL=<0x0002> <MR>
(0x0008,0x0070) LO Manufacturer VR=<LO> VL=<0x0012> <GE MEDICAL SYSTEMS>
(0x0008,0x0080) LO Institution Name VR=<LO> VL=<0x001c> <THOMAS JEFF UNIVHOSPITAL MRI>
(0x0008,0x0090) PN Referring Physician's Name VR=<PN> VL=<0x0004> <HUME>
(0x0008,0x1010) SH Station Name VR=<SH> VL=<0x0008> <FOR.IC0 >
(0x0008,0x1030) LO Study Description VR=<LO> VL=<0x0004> <KNEE>
(0x0008,0x103e) LO Series Description VR=<LO> VL=<0x0006> <COR T2>
(0x0008,0x1060) PN Name of Physician(s) Reading Study VR=<PN> VL=<0x0004> <BODY>
(0x0008,0x1070) PN Operator's Name VR=<PN> VL=<0x0002> <RB>
(0x0008,0x1090) LO Manufacturer's Model Name VR=<LO> VL=<0x000e> <GENESIS_SIGNA >
(0x0010,0x0010) PN Patient's Name VR=<PN> VL=<0x000c> <* GRX KNEE *>
(0x0010,0x0020) LO Patient's ID VR=<LO> VL=<0x0006> <RSNA2 >
(0x0010,0x0030) DA Patient's Birth Date VR=<DA> VL=<0x0000> []
(0x0010,0x0040) CS Patient's Sex VR=<CS> VL=<0x0002> <M >
(0x0010,0x1010) AS Patient's Age VR=<AS> VL=<0x0004> <034Y>
(0x0010,0x1030) DS Patient's Weight VR=<DS> VL=<0x000a> < 90.718000>
(0x0010,0x21b0) LT Additional Patient History VR=<LT> VL=<0x0008> <R/O TEAR>
(0x0018,0x0020) CS Scanning Sequence VR=<CS> VL=<0x0002> <SE>
(0x0018,0x0021) CS Sequence Variant VR=<CS> VL=<0x0004> <OSP >
(0x0018,0x0022) CS Scan Options VR=<CS> VL=<0x0004> <NPW >
(0x0018,0x0023) CS MR Acquisition Type VR=<CS> VL=<0x0002> <2D>
(0x0018,0x0025) CS Angio Flag VR=<CS> VL=<0x0002> <N >
(0x0018,0x0050) DS Slice Thickness VR=<DS> VL=<0x0008> <5.000000>
(0x0018,0x0080) DS Repetition Time VR=<DS> VL=<0x000c> < 2000.000000>
(0x0018,0x0081) DS Echo Time VR=<DS> VL=<0x000a> < 20.000000>
(0x0018,0x0082) DS Inversion Time VR=<DS> VL=<0x0008> <0.000000>
(0x0018,0x0083) DS Number of Averages VR=<DS> VL=<0x0008> <0.500000>
(0x0018,0x0084) DS Imaging Frequency VR=<DS> VL=<0x0010> < 638746840.00000>
(0x0018,0x0085) SH Imaged Nucleus VR=<SH> VL=<0x0002> <H1>
(0x0018,0x0086) IS Echo Number(s) VR=<IS> VL=<0x0002> < 1>
(0x0018,0x0087) DS Magnetic Field Strength VR=<DS> VL=<0x0006> <15000 >
(0x0018,0x0088) DS Spacing Between Slices VR=<DS> VL=<0x0008> <6.000000>
(0x0018,0x0091) IS Echo Train Length VR=<IS> VL=<0x0002> <0 >
(0x0018,0x0093) DS Percent Sampling VR=<DS> VL=<0x000a> < 53.125000>
(0x0018,0x0094) DS Percent Phase Field of View VR=<DS> VL=<0x000a> <100.000000>
(0x0018,0x1088) IS Heart Rate VR=<IS> VL=<0x0002> <0 >
(0x0018,0x1090) IS Cardiac Number of Images VR=<IS> VL=<0x0002> <0 >
(0x0018,0x1094) IS Trigger Window VR=<IS> VL=<0x0002> <10>
(0x0018,0x1100) DS Reconstruction Diameter VR=<DS> VL=<0x000a> <140.000000>
(0x0018,0x1314) DS Flip Angle VR=<DS> VL=<0x0002> <0 >
(0x0018,0x1315) CS Variable Flip Angle Flag VR=<CS> VL=<0x0002> <N >
(0x0018,0x1316) DS SAR VR=<DS> VL=<0x0008> <0.052993>
(0x0018,0x5100) CS Patient Position VR=<CS> VL=<0x0004> <FFS >
(0x0020,0x000d) UI Study Instance UID VR=<UI> VL=<0x0028>
<1.2.840.113619.2.1.2.139348932.602501178>
(0x0020,0x000e) UI Series Instance UID VR=<UI> VL=<0x002a>
<1.2.840.113619.2.1.2.596272627.1.602501541>
(0x0020,0x0010) SH Study ID VR=<SH> VL=<0x0002> <2 >
(0x0020,0x0011) IS Series Number VR=<IS> VL=<0x0002> < 1>
(0x0020,0x0012) IS Acquisition Number VR=<IS> VL=<0x0002> < 0>
(0x0020,0x0013) IS Image Number VR=<IS> VL=<0x0002> < 1>
(0x0020,0x0032) DS Image Position (Patient) VR=<DS> VL=<0x0020>
<-70.000000\ 18.000000\ 75.000000>
(0x0020,0x0037) DS Image Orientation (Patient) VR=<DS> VL=<0x0038>
< 1.000000\0.000000\0.000000\0.000000\0.000000\ -1.000000>
(0x0020,0x0052) UI Frame of Reference UID VR=<UI> VL=<0x002c>
<1.2.840.113619.2.1.2.596272627.1.602501541.0>
(0x0020,0x0060) CS Laterality VR=<CS> VL=<0x0000> []
(0x0020,0x0110) DS Temporal Resolution VR=<DS> VL=<0x000a> <1120403456>
(0x0020,0x1040) LO Position Reference Indicator VR=<LO> VL=<0x0002> <KN>
(0x0020,0x1041) DS Slice Location VR=<DS> VL=<0x000e> <-18.0000000000>
(0x0028,0x0002) US Samples per Pixel VR=<US> VL=<0x0002> [0x01]
(0x0028,0x0004) CS Photometric Interpretation VR=<CS> VL=<0x000c> <MONOCHROME2 >
(0x0028,0x0010) US Rows VR=<US> VL=<0x0002> [0x100]
(0x0028,0x0011) US Columns VR=<US> VL=<0x0002> [0x100]
(0x0028,0x0030) DS Pixel Spacing VR=<DS> VL=<0x0012> < 0.546875\0.546875>
(0x0028,0x0100) US Bits Allocated VR=<US> VL=<0x0002> [0x10]
(0x0028,0x0101) US Bits Stored VR=<US> VL=<0x0002> [0x10]
(0x0028,0x0102) US High Bit VR=<US> VL=<0x0002> [0x0f]
(0x0028,0x0103) US Pixel Representation VR=<US> VL=<0x0002> [0x01]
(0x0028,0x0120) XS Pixel Padding Value VR=<SS> VL=<0x0002> [0x00]
(0x7fe0,0x0010) OX Pixel Data VR=<OW> VL=<0x20000> []

## Slide 41: Encoding – Sequences
Sequence (SQ) VR allows “nesting” of lists
Data Element with an SQ VR
  no “value” field per se
  may have a fixed or undefined length
  zero or more Sequence Items
  Sequence Delimiter tag (if was undefined length)
Sequence Item
  Item tag (may have a fixed or undefined length)
  list of (sorted, unique) data elements
  Item Delimiter tag (if was undefined length)

## Slide 43: Dump Tool Output for Sequence
%item
(0x0040,0xa010) Relationship Type <HAS OBS CONTEXT>
(0x0040,0xa040) Value Type <PNAME >
(0x0040,0xa043) Concept Name Code Sequence
%item
(0x0008,0x0100) Code Value <000555>
(0x0008,0x0102) Coding Scheme Designator <LNdemo>
(0x0008,0x0104) Code Meaning <Recording 0bserver>
%enditem
%endseq
(0x0040,0xa123) Person Name <Smith^John^^Dr^ >
%enditem

## Slide 44: Actual Binary Encoding
...
(0x0028,0x0002) Samples per Pixel VR=<US> VL=<0x0002> [0x0001]
(0x0028,0x0004) Photometric Interpretation VR=<CS> VL=<0x000c> <MONOCHROME2 >
(0x0028,0x0008) Number of Frames VR=<IS> VL=<0x0004> <124 >
(0x0028,0x0010) Rows VR=<US> VL=<0x0002> [0x0100]
(0x0028,0x0011) Columns VR=<US> VL=<0x0002> [0x0100]
(0x0028,0x0100) Bits Allocated VR=<US> VL=<0x0002> [0x0010]
(0x0028,0x0101) Bits Stored VR=<US> VL=<0x0002> [0x0010]
(0x0028,0x0102) High Bit VR=<US> VL=<0x0002> [0x000f]
...
...
00000560 .. .. .. .. .. .. 28 00 02 00 55 53 02 00 01 00 |..........US....|
00000570 28 00 04 00 43 53 0c 00 4d 4f 4e 4f 43 48 52 4f |(...CS..MONOCHRO|
00000580 4d 45 32 20 28 00 08 00 49 53 04 00 31 32 34 20 |ME2 (...IS..124 |
00000590 28 00 10 00 55 53 02 00 00 01 28 00 11 00 55 53 |(...US....(...US|
000005a0 02 00 00 01 28 00 00 01 55 53 02 00 10 00 28 00 |....(...US....(.|
000005b0 01 01 55 53 02 00 10 00 28 00 02 01 55 53 02 00 |..US....(...US..|
...
*names shown in test tool dump are not actually encoded

## Slide 45: You Need a Library or Toolkit
You can hand-code reading & writing
But:
  legal variations in encoding (Transfer Syntax)
  sequence handling is non-trivial
  need a “data dictionary” to understand tags
  need an “IOD” library to map flat list of data elements as encoded to/from IODs/Modules (or “classes”)
  need utilities for dumping, testing, editing, etc.
Not to mention files, services, protocols!
Fortunately, are available for all platforms

## Slide 46: Private Data Elements
Odd group numbers are all private
(gggg,00xx) is a private creator string
(gggg,xxyy) is the block defined by that creator
(0019,0010) “David’s Stuff”
(0019,0011) “Harry’s Stuff”
(0019,1001) 1st of david’s private data elements
(0019,1002) 2nd of david’s private data elements
…
(0019,1101) 1st of harry’s private data elements
(0019,1102) 2nd of harry’s private data elements
…

## Slide 47: DICOM – Is it a “File Format”?
Yes and no
  since 1995 – a formal “file format” has been defined (PS3.10 aka. “part 10”)
  but, it applies only to files on specified media (e.g., CD) with DICOMDIR (directory) (“File Set”)
  strictly speaking, no standard conformance requirements for files “by themselves”
In practice
  esp. for research and testing, “DICOM files” are often stored and exchanged informally
  even old ACR-NEMA data sets were stored in files and interchanged

## Slide 48: File Format v. Data Set
Data Set
  “an instance of a real world Information Object”
  “constructed of Data Elements”
  is what is exchanged on network using the C-STORE “command”
File Format
  “a means to encapsulate in a file the Data Set”
  “byte stream of the Data Set is placed into the file after the DICOM File Meta Information”
  “each file contains a single SOP Instance”

## Slide 49
Data Set Data Elements
Data Set

## Slide 50
Data Set Data Elements
Pixel Data Element
Data Set

## Slide 51
Data Set Data Elements
Pixel Data Element
Data Set Data Elements
Data Set

## Slide 52
Data Set Data Elements
Pixel Data Element
Data Set Data Elements
Data Set
“header”

## Slide 53
Data Set Data Elements
Pixel Data Element
Data Set Data Elements
Data Set
“header”
offset to start of pixel data not fixed length

## Slide 54
128 byte preamble (ignored)
“DICM” prefix (magic #)
File Meta Information Data Elements
Data Set Data Elements
Pixel Data Element
Data Set Data Elements
File Meta Information
Data Set

## Slide 55
128 byte preamble (ignored)
“DICM” prefix (magic #)
File Meta Information Data Elements
Data Set Data Elements
Pixel Data Element
Data Set Data Elements
File Meta Information
Data Set
“header”
“header”

## Slide 56
128 byte preamble (ignored)
“DICM” prefix (magic #)
File Meta Information Data Elements
Data Set Data Elements
Pixel Data Element
Data Set Data Elements
File Meta Information
Data Set
DICOM File

## Slide 57
128 byte preamble (ignored)
“DICM” prefix (magic #)
File Meta Information Data Elements
Data Set Data Elements
Pixel Data Element
Data Set Data Elements
File Meta Information
Data Set
Network

## Slide 58: Why Meta Information Header?
Make the “file” recognizable as a DICOM file
  if just Data Set stored, would have to “guess”
Describe the encoding of the Data Set
  the “Transfer Syntax”
  little or big endian? compressed?
  negotiated on network; needs to to be explicitly specified for file
Describe contents of the Data Set
  SOP Class and SOP Instance UIDs
  allows decision to use or ignore before decoding Data Set

## Slide 59: Transfer Syntax
Term inherited from ISO/OSI
“a set of encoding rules”
Uncompressed
  Default: little endian implicit VR (no “type” is sent)
  Little endian explicit VR (“type” is sent)
  Big endian explicit VR (byte order affected speed)
Compressed
  Entire Data Set: Deflate (gzip-like) (rarely used)
  Pixel Data: JPEG (8,12,DCT,lossless), JPEG-LS, JPEG 2000, MPEG2, MPEG4, RLE

## Slide 60: DICOM Services - PACS
Modality
Archive
Modality
Modality
Modality
PACS +/- RIS
Manager
Workstations
DICOM
DICOM

## Slide 61: Original DICOM Services
Verification
Storage
Query/Retrieve
Study Content Notification (*)
Detached Management (*)
  patient/visit/study/results/interpretation
Print Management (* = not widely used)

## Slide 62: Newer DICOM Services
Media Storage
Storage Commitment
(Modality) Worklist Management
(Modality) Performed Procedure Step
…
Display Function Standard
Structured Reporting

## Slide 63: Query/Retrieve Classes
Hierarchical or relational (*)
Models
  Patient/Study/Series/Image
  Study (including Patient)/Series/Image
  Patient/Study only (*)
Services:
  Find, Move, Get (*) (* = rarely used)

## Slide 64: Print Service Classes
Basic vs. Referenced
  Images included or referenced
  Basic is far more common
Grayscale vs. Color
  grayscale for traditional laser images
  color for dye sub, etc.
  CR, CT, MR vs. NM, US

## Slide 65: Modality/IS Related Services
Modality Worklist (MWL)
  modality queries IS for list of work
    patient name, id etc.
    attributes to identify images of study in stored output (UIDs)
Modality Performed Procedure Step (MPPS)
  modality reports what was done
Storage Commitment
  modality asks archive/workstation to take responsibility for subsequent storage of images
Basis for IHE Scheduled Workflow Profile

## Slide 66: IHE Scheduled Workflow

## Slide 67: IHE Scheduled Workflow
HL7V2
HL7V2
DICOM MWL
DICOM STORE MPPS COMMIT
DICOM STORE Q/R

## Slide 68: Display Function Standard
Inconsistencies of grayscale display
  monitor vs. workstation
  display vs. printer
  between printers (vendors, calibration)
Standard curve
  perceptually linearized (Barten)
  input step is Just Noticeable Difference
  output is luminance, optical density
  calibration of display device

## Slide 69: Example of DICOM Services
CT scanner
  acquire; archive and print over the network
Workstation
  3D reconstruction of CT and MR images
  store results on network and media
  print over the network
Archive
Printer (laser imager)

## Slide 70: DICOM Services – No PACS
CT Modality
Laser Printer
Shared Archive
Workstation
Store
Store
Store
Print
Print
Q/R
Q/R

## Slide 71: Conformance Statement
DICOM Conformance Statement
  defines services
    including modality-specific storage classes
  defines roles (user or provider)
  defines other specific conformance issues
    Transfer Syntax (encoding) - there is a default
    limitations
    configuration
    physical network
  required by the standard, in a standard format
  a public document, usually shared on web site

## Slide 72: Conformance Statement Example CT Scanner Specifications - SCU Role
“This Application Entity provides Standard Conformance to the following DICOM v3.0 SOP Classes as an SCU:”

## Slide 73: Conformance Statement Example CT Scanner Specifications - SCP Role
“This Application Entity provides Standard Conformance to the following DICOM v3.0 SOP Classes as an SCP:”

## Slide 74: Integration using DICOM
Define primary functions
Define boundaries between functions
Identify equipment with functions
Define primary services
Define support services
Identify DICOM service/role per device
Match Conformance Statements
Don’t assume “plug and play”
Define expected functionality a priori
Just specifying “DICOM” is meaningless

## Slide 75: Integration using DICOM
Delineation of Responsibility
One source
  vendor takes responsibility
Multiple sources in one purchase
  integrator (user ?) takes responsibility
Multiple sources, multiple purchases
  integrator ? user ? new supplier ?
Upgrades
  upgrade supplier takes responsibility
  evolution v. revolution (data/image migration)

## Slide 76: More DICOM Specifics
Information Object Definitions
Structures and Encoding
Media Format
Service Classes
Upper Layer Protocol (DUL)
Association Establishment (ACSE)
Message Service Element (DIMSE)

## Slide 77: Service Classes
Verification
Storage
Query/Retrieve
Print Management
Storage Commitment
Basic Worklist Management.
Modality Performed Procedure Step
…

## Slide 78: Service Classes
Concept of user (SCU), provider (SCP)
  like “client” and “server”, or “source” and “sink”
Message services used (DIMSE)
  e.g., Storage Service uses C-STORE message
  e.g., Query Service uses C-FIND message
Service-Object Pair (SOP) Classes
  combine a service with an object (IOD)
  e.g., CT Image Storage SOP Class
    Storage Service Class
    CT Image IOD

## Slide 79: DICOM Application Layer

## Slide 80: DICOM Upper Layer Protocol

## Slide 81: DICOM Upper Layer Protocol

## Slide 82: DICOM Upper Layer Protocol
TCP/IP over usual physical layers
  may be secured using TLS (or VPN)
Presentation Data Service
  P-DATA-TF
Association Control Service Element (ACSE)
  A-ASSOCIATE-RQ,-AC,-RJ
  A-RELEASE-RQ,-RSP
  A-ABORT

## Slide 83: DUL – Addressing
Who am I and who am I talking to ?
Application Entity Titles (AETs)
Calling AET
  tells other end who (what device) I am
  other end can filter on this (access control)
  logging/audit trail
Called AET
  who (what device) I want to talk
  need mapping to Presentation Address
    IP address (or hostname via table/NIS/DNS)
    Port: 11112 or 104 (though often others used)

## Slide 84: Association Establishment
Association Control Service Element (ACSE)
Key feature is negotiation
  Presentation Context
    Abstract Syntax (SOP Class)
    Transfer Syntax (encoding, compression)
  Other parameters
    Maximum PDU length
    Implementation Class UID
    Asynchronous Operations
    SCP/SCU Role selection

## Slide 85: DUL – Association Request

## Slide 86: Presentation Context

## Slide 87: A-ASSOCIATE-RQ/-AC PDU

## Slide 88: Message Service Elements
DIMSE-C
  C-STORE
  C-GET
  C-MOVE
  C-FIND
  C-ECHO
  C-CANCEL
DIMSE-N
  N-EVENT-REPORT
  N-GET
  N-SET
  N-ACTION
  N-CREATE
  N-DELETE

## Slide 89: DIMSE Service Primitives

## Slide 90: DIMSE Protocol - Commands

## Slide 91: P-DATA-TF PDU

## Slide 92: DUL and DIMSE Example
SCU
SCP
Establish TCP/IP Connection
A-ASSOCIATE-RQ
to IP/port deduced from AET
offer CT Image Storage, default TS
A-ASSOCIATE-AC
accept CT Image Storage, default TS
P-DATA-TF
C-STORE-RQ Msg: Command
P-DATA-TF
C-STORE-RQ Msg: Data Set
P-DATA-TF
C-STORE-RSP Msg: Command
A-RELEASE-RQ
A-RELEASE-RSP
Close TCP/IP Connection

## Slide 93: DUL and DIMSE Example
SCU
SCP
Establish TCP/IP Connection
A-ASSOCIATE-RQ
to IP/port deduced from AET
offer CT Image Storage, TS1, TS2
A-ASSOCIATE-AC
accept CT Image Storage, TS1
P-DATA-TF
C-STORE-RQ Msg: Command
Loop
P-DATA-TF
C-STORE-RQ Msg: Data Set
P-DATA-TF
C-STORE-RSP Msg: Command
A-RELEASE-RQ
A-RELEASE-RSP
Close TCP/IP Connection

## Slide 94: DUL and DIMSE Example
Loop
SCU
SCP
Establish TCP Connection
A-ASSOCIATE-RQ
to IP/port deduced from AET
offer CT Image Storage, TS1, TS2
A-ASSOCIATE-AC
accept CT Image Storage, TS1
P-DATA-TF
C-STORE-RQ Msg: Command
Loop
P-DATA-TF
C-STORE-RQ Msg: Data Set
P-DATA-TF
C-STORE-RSP Msg: Command
A-RELEASE-RQ
A-RELEASE-RSP
Close TCP Connection

## Slide 95: Summary
Brief history
File formats and Storage Services
Data Set encoding principles
Transfer Syntaxes
Information Object Definitions
Storage SOP Classes
Other DICOM services than storage
Conformance statements
Association Negotiation
DICOM Message Service Elements (DIMSE)
DICOM Upper Layer Protocol (DUL)